## Slide 1
Important Features of a Pipeline for Processing Diffusion Weighted Images
Carlo Pierpaoli
Quantitative Medical Imaging Section
NIBIB

## Slide 2: Quantitative diffusion MRI requires multiple coregistered images
8
200
400
650
b (s/mm2)
Diffusion coefficient map
S = S0 e-bD
Notes: Typically, we take the log of the equation and fit the linear form

## Slide 3
Assessment diffusion in anisotropic tissues requires multiple images with non-collinear gradient directions
Radiology. 1996 Dec;201(3):637-48

## Slide 4
Importing data, check if images are sorted correctly
Initial quality control by visual inspection
Production of “corrected” raw images
- EPI distortion correction
- Motion and eddy-current distortion correction
- Registration to standard template
- Rotation of b-matrices
- Generate interpolation-corrected weighting
coefficients
Tensor computation (correct weighting)
Analysis of residuals of tensor fitting, inspection maps
Robust tensor estimation
Computation of tensor derived variables or HARDI analysis
Tasks to address in a DWI processing pipeline

## Slide 5
Importing data correctly
Gradient conventions are not standardized across vendors
Possibility of wrong information being reported in the DICOM header
Handling of oblique DWI acquisitions is inconsistent across vendors and even across software releases from the same vendor

## Slide 6
DEC map (absolute value scheme) may look correct

## Slide 7
Underlying vector field is incorrect

## Slide 8
Magnetic Resonance in Medicine 42:526–540 (1999)

## Slide 9
USE no symmetry DEC map to verify correct eigenvector
orientation

## Slide 10
Subject motion

## Slide 11
Reference
Image
Shift
Bo along Z
Gradient along X
Gradient along Y
Shear
Stretch /
Compress
Eddy Current Induced Distortions
Courtesy of Derek Jones who in turn thanks Joe Zhou

## Slide 12
Eddy Current Induced Distortions

## Slide 13
Effects of image mis-registration on tensor-derived parameters Artifactually high anisotropy…
FA
DEC
Dxy
Dxz
Dyz

## Slide 14
DEC map computed from data corrupted by severe motion
DEC map computed after motion correction
Effects of image mis-registration on tensor-derived parameters ….and/or artifactually low anisotropy

## Slide 15
Remedies for misregistration artifacts
Eddy current distortion:
acquisition (e.g. twice refocused spin echo, Reese 2003)
field maps (Jezzard 1998)
image registration (Haselgrove,1996; Bastin 1999; Horsfield, 1999; Andersson, 2002; Rohde, 2004)
opposite diffusion gradient polarity (Bodammer 2004,
Andersson, 2016)
Subject motion:
image registration (Andersson, 2002; Rohde, 2004)

## Slide 16
Desirable features in registration algorithms for diffusion MRI
Motion correction, eddy current correction, and registration to a standard template done with only one interpolation step

## Slide 17
Desirable features in registration algorithms for diffusion MRI
Motion correction, eddy current correction, and registration to a standard template done with only one interpolation step
Correct deformation model

## Slide 18: Eddy distortion correction head to head comparison
AFFINE registration
Affine + Quadratic TORTOISE
Notes: Not all algorithms are created equally. Affine versus higher order term (reference ASB)

## Slide 19
Desirable features in registration algorithms for diffusion MRI
Motion correction, eddy current correction, and registration to a standard template done with only one interpolation step
Correct deformation model
- Parsimonious deformation model

## Slide 20
Desirable features in registration algorithms for diffusion MRI
Motion correction, eddy current correction, and registration to a standard template done with only one interpolation step
Correct deformation model
- Parsimonious deformation model
- No additional data needed (e.g. no phantom calibration)

## Slide 21
Desirable features in registration algorithms for diffusion MRI
Motion correction, eddy current correction, and registration to a standard template done with only one interpolation step
Correct deformation model
- Parsimonious deformation model
- No additional data needed (e.g. no phantom calibration)
- Appropriate reorientation of b-matrices

## Slide 22
Desirable features in registration algorithms for diffusion MRI
Motion correction, eddy current correction, and registration to a standard template done with only one interpolation step
Correct deformation model
- Parsimonious deformation model
- No additional data needed (e.g. no phantom calibration)
- Appropriate reorientation of b-matrices
- co-registration of images with very different contrast

## Slide 23
T2WI
DWI
Registration based on mutual information
DWI
Registration based on correlation

## Slide 24
Example of correction of eddy-current and motion artifacts (according to Rohde et al. MRM 2004)

## Slide 25
Correction of eddy-current and motion artifacts according to Rohde et al. MRM 2004
DEC
FA
Computed from unregistered raw images
Computed from registered raw images

## Slide 26
Principal Component Analysis of signals in a DWI data set
From Rohde et al, MRM 2004
Principal components are displayed in order of decreasing variance from top-left
No motion
Notes: The n volumes of a dti dataset can be seen as a random vector of n components and m elements where m is the number of voxels in the imaged volume.
We compute a new set of n volumes
PCA reference: Anderson TW 1957

## Slide 27
Principal components are displayed in order of decreasing variance from top-left
In presence of
uncorrected
motion
Principal Component Analysis of signals in a DWI data set
From Rohde et al, MRM 2004
Notes: The n volumes of a dti dataset can be seen as a random vector of n components and m elements where m is the number of voxels in the imaged volume.
We compute a new set of n volumes
PCA reference: Anderson TW 1957

## Slide 28
B = 30 s/mm2
B = 900 s/mm2
B = 9000 s/mm2
Registration of high b-value images: Challenging…
.
…but interesting approaches have been proposed
(e.g. J.Andersson, S Sotiropoulos NEUROIMAGE 2016, 125 1063-78)

## Slide 29
EPI distortion

## Slide 30
T1 Target
before correction
after correction
EPI distortion

## Slide 31
No EPI Correction
With EPI Correction (Wu et al, MICCAI 2008
Mean FA
SD of FA
SD of FA
Voxelwise analysis of tensor derived quantities from
Living Phantom data (21 scans from 5 sites)
Notes: Living Phantom standard deviation across N=21 scans from 5 sites (BOS = 6, CIN = 4, PHI = 5, UCLA = 3, STL=3). These were processed with the individual target T2 which were all rigidly co-registered. Mean FA is from the EPI corrected data. All data were corrected for Eddy and Motion.

## Slide 34
Native EPI distortion
EPI distortion corrected

## Slide 35: EPI distortion correction strategies
Field-mapping (Jezzard et al, 1995).
  Requires estimated B0 inhomogeneity maps from two scans with different echo times.
  Accurate estimation of the B0 map may be challenging.
  Can not properly correct regions of severe signal pile-up.
EPI distortion correction strategies can be broadly categorized into three classes:
  Field-mapping.
  Image registration.
  Reverse phase encoding (Blip-up blip-down).

## Slide 36
Elastic image registration based approaches (Kybic et al, 2000, Wu et al. 2008).
  Requires a distortion-free anatomical image.
  Can not properly correct regions of severe signal pile-up.
Constrained elastic registration

## Slide 37
Reversed phase encoding (blip-up blip-down) correction (Chang et al. 1992, Bowtell, 1994, Andersson et al., 2003, Holland et al., 2010.)
It exploits the symmetry of distortions in two identical images acquired with opposite polarity of phase encoding
Acquired with RL phase encoding
Acquired with LR phase encoding

## Slide 38
  Only strategy that can correct for signal pile-ups and expansions, as well as geometrical distortions.
  More robust with larger deformations than standard elastic registration approaches.
  May need longer scan time
Blip-up blip-down data
Corrected data

## Slide 40
Important features of DR-BUDDI (Diffeomorphic Registration for Blip-Up blip-Down Diffusion Imaging)
Can use a structural image as additional information for registration (fat suppressed T2-WI)
Performs blip-up blip down correction also on DWIs
All corrections are applied to the original images with one interpolation step
Diffeomorphic
Symmetric
Large deformation
Constrained

## Slide 41: DR-BUDDI pipeline

## Slide 42: DR-BUDDI Results
Correction of highly distorted CONNECTOME data
Notes: 3/27/19
42

## Slide 43
Cardiac Pulsation
Cardiac pulsation has been identified as an important source of variability in fMRI.
Previous studies have shown some benefits of cardiac gating in DWI and apparent diffusion coefficient measurements . (Chien et al. 1990, Turner et al. 1990, Conturo et al. 1995, Skare et al. 2001)
However, the effect of cardiac pulsation on DTI measurements has not been extensively investigated and cardiac gating is rarely used in clinical DT-MRI.

## Slide 44
Trace(D)
Kruskal-Wallis test
Statistically significant voxels in red
Bonferroni correction
C Pierpaoli – Artifacts in Diffusion MRI in Diffusion MRI: Theory, Methods and Applications, 2010

## Slide 45
FA
Kruskal-Wallis test
Statistically significant voxels in red
Bonferroni correction
C Pierpaoli – Artifacts in Diffusion MRI in Diffusion MRI: Theory, Methods and Applications, 2010

## Slide 47
Cardiac pulsation significantly affects, Trace(D), anisotropy measures, and apparent fiber orientation in several brain regions
Most of the variability is produced by data acquired from 100 to 250 ms after the QRS

## Slide 48
Possible solutions
1) Cardiac gating
- in product sequences: not widely used, not optimized, sometimes buggy
- believed to increase scan time (somewhat incorrectly because often the rate limiting factor is gradient duty cycle).
2) Retrospective gating
3) Robust tensor fitting

## Slide 52
Goodness of fit criterion
satisfied?
Final nonlinear LS fit
Convergence?
Reset weights
based on the residuals
Nonlinear robust fit
Tensor Maps
Iterative re-weighting process
Outlier identification
and exclusion
Nonlinear LS fit
Yes
No
No
Yes
RESTORE:
Robust
EStimation of
Tensors by
Outlier
REjection
Chang et al,
MRM 2005

## Slide 53
RESTORE performance. Monte carlo simulation

## Slide 54: Automatic exclusion of “bad images”

## Slide 55: After registration the “bad” image is distributed across several slices

## Slide 56: One “bad” image out of 28 DWIs can significantly corrupt tensor -derived quantities
Conventional tensor fitting
Robust tensor fitting using RESTORE

## Slide 57
Robust tensor fitting: Selective (voxel-based) exclusion of corrupted data

## Slide 58
Robust tensor fitting: Selective (voxel-based) exclusion of corrupted data

## Slide 59
Robust tensor fitting: Selective (voxel-based) exclusion of corrupted data

## Slide 60
RESTORE performance. Cardiac Pulsation
T2wi
Outlier map
Outlier map
Diastolic data
Diastolic + systolic data

## Slide 61: Outlier Rejection Probability Map Lindsay Walker et al Neuroimage 2011

## Slide 62
Biological Psychiatry - 2013
Importance of outlier and artifact detection in clinical imaging

## Slide 63
B) Clusters of significant differences in MD (AUT > TYP)
A) Clusters of significant differences in FA (TYP > AUT)
white matter skeleton voxels
statistically significant voxels
Notes: Figure 1

## Slide 64
Percentage of Outlier Data Points in a single subject example
up to 50% outliers detected
~3% outliers detected
Notes: Supplemental figure S5

## Slide 65
-7%
7%
Subtraction of Percentage of Outliers – AUT minus TYP
Notes: Figure 4 – Difference in distribution of outliers

## Slide 66
A good and somewhat forgotten “diagnostic” tool to help detecting artifacts in diffusion MRI
Map of reduced 2 (goodness of fit) from tensor fitting (Papadakis JMR 2003)

## Slide 67
Characterization of normal brain and brain tumor pathology by chisquares parameter maps of diffusion-weighted image data
Maier SE, Mamata H, Mulkern RV
EUROPEAN JOURNAL OF RADIOLOGY
45 (3): 199-207 MAR 2003
A measure of curve fitting error for noise filtering diffusion tensor MRI data
Papadakis NG, Martin KM, Wilkinson ID, Huang CLH
JOURNAL OF MAGNETIC RESONANCE
164 (1): 1-9 SEP 2003
Relevant references

## Slide 68
Ref: Bevington PR, Robinson DK. 3rd Ed. 2003.
2 = S2 / i
iexpected variance

## Slide 69
If thermal noise is the only source of variability and the tensor model fits the data perfectly …
….The 2 map should be homogeneous with an average value of 1
2 = measured variance (S2) /
expected variance (i

## Slide 70
2 maps. Healthy human brain, cardiac gating, motion correction
Grayscale: 0 black, 5 brightest

## Slide 71
Examples of the utility of 2 for detecting artifacts in diffusion MRI

## Slide 72
Examples of the utility of 2 for detecting artifacts in diffusion MRI

## Slide 73
Cause of the artifact: imperfect fat suppression

## Slide 74
DEC
FA
2

## Slide 75
Residual histogram in TORTOISE

## Slide 76
Residual histogram. What problem are we detecting?

## Slide 77
Processing pipeline should enable checking for artifacts like this
….

## Slide 78
First DWI
Last DWI after 20 min of scanning

## Slide 79
First DWI
Last DWI after 20 min of scanning
B0 Frequency shift caused by gradient heating

## Slide 80
Tasks to address in a DTI processing pipeline

## Slide 81
Importing data, check if images are sorted correctly
Initial quality control by visual inspection
Tasks to address in a DTI processing pipeline

## Slide 82
Importing data, check if images are sorted correctly
Initial quality control by visual inspection
Production of “corrected” raw images
- EPI distortion correction
- Motion and eddy-current distortion correction
- Registration to standard template
- Rotation of b-matrices
- Generate interpolation-corrected weighting
coefficients
Tasks to address in a DTI processing pipeline

## Slide 83
Importing data, check if images are sorted correctly
Initial quality control by visual inspection
Production of “corrected” raw images
- EPI distortion correction
- Motion and eddy-current distortion correction
- Registration to standard template
- Rotation of b-matrices
- Generate interpolation-corrected weighting
coefficients
Tensor computation (correct weighting)
Tasks to address in a DTI processing pipeline

## Slide 84
Importing data, check if images are sorted correctly
Initial quality control by visual inspection
Production of “corrected” raw images
- EPI distortion correction
- Motion and eddy-current distortion correction
- Registration to standard template
- Rotation of b-matrices
- Generate interpolation-corrected weighting
coefficients
Tensor computation (correct weighting)
Analysis of residuals of tensor fitting, inspection maps
Tasks to address in a DTI processing pipeline

## Slide 85
Importing data, check if images are sorted correctly
Initial quality control by visual inspection
Production of “corrected” raw images
- EPI distortion correction
- Motion and eddy-current distortion correction
- Registration to standard template
- Rotation of b-matrices
- Generate interpolation-corrected weighting
coefficients
Tensor computation (correct weighting)
Analysis of residuals of tensor fitting, inspection maps
Robust tensor estimation
Tasks to address in a DTI processing pipeline

## Slide 86
Importing data, check if images are sorted correctly
Initial quality control by visual inspection
Production of “corrected” raw images
- EPI distortion correction
- Motion and eddy-current distortion correction
- Registration to standard template
- Rotation of b-matrices
- Generate interpolation-corrected weighting
coefficients
Tensor computation (correct weighting)
Analysis of residuals of tensor fitting, inspection maps
Robust tensor estimation
Computation of tensor derived variables
Tasks to address in a DTI processing pipeline

## Slide 87
Importing data, check if images are sorted correctly
Initial quality control by visual inspection
Production of “corrected” raw images
- EPI distortion correction
- Motion and eddy-current distortion correction
- Registration to standard template
- Rotation of b-matrices
- Generate interpolation-corrected weighting
coefficients
Tensor computation (correct weighting)
Analysis of residuals of tensor fitting, inspection maps
Robust tensor estimation
Computation of tensor derived variables and/or HARDI analysis
Tasks to address in a DTI processing pipeline